IEEE P802.11
Wireless LANs

| COAT related Comment Resolution |  |  |  |  |
| --- | --- | --- | --- | --- |
| Date: 2008-01-14 |  |  |  |  |
| Author(s): |  |  |  |  |
| Name | Affiliation | Address | Phone | email |
| Dalton Victor | Broadcom Corporation | 190 Mathilda Place Sunnyvale, CA 94086 | 408-922-5824 | dvictor@broadcom.com |
| Mark Kobayashi | Broadcom Corporation | 190 Mathilda Place Sunnyvale, CA 94086 | 408-543-3318 | mmk@broadcom.com |

Abstract

- This draft text addresses comments on section 5.3.

- The number of CIDs resolved by this document is 30.

| 173 | Foegelle, Michael | 5.3 | 18.44 | T | This section attempts to define an environment and procedure for accurate OTA testing, but falls far short of the mark. The fundamentally flawed assumption is that placing the DUT antenna at the exact point of the reference antenna somehow guarantees that the applied calibration will produce accurate results. This is far from the case, since the actual radiation pattern of the DUT (not just the antenna, but the DUT body, circuitry, cables, etc. all of which form part of the radiating object) is not known. The procedure again makes the fundamentally flawed assumption that just because the results are repeatable, that they are accurate. If the wrong thing is done over and over with the same results, it doesn't make those results right. While ideally a DUT would have a perfectly symmetrical omnidirectional pattern oriented in the horizontal plane, that is rarely the case. It can have significant peaks and nulls. In this case, if the DUT happens to be placed such that it has a peak pointed towards the measurement antenna, it will appear to perform better than another DUT that may have better overall performance but have a null pointed in that direction when placed in the same orientation. It will even perform differently if oriented differently about its own presumed polarization axis. Thus, this method as written can be used to obtain any desired result when comparing the performance of two products. What good is that? It's not traceable and it's not acceptable. The only way to guarantee traceable quantitative results that can be compared is to measure the entire radiation pattern of the active device and use that information to provide a total performance metric, rather than one from only one chosen direction, which may or may not be any good. | Implement the industry standard TRP/TIS measurements that perform the desired active pattern measurements to determine the true performance of the device. Implement the procedures in 11-06 0131r1. |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 140 | Foegelle, Michael | 5.3.1 | 18.48 | T | The claim of yielding accurate results is unfounded based on the existing description of the environment and associated test procedures. Repeatability does not imply accuracy, and it is easy to prove that the assumptions made about the behavior of the DUT can completely invalidate the results even if the path loss has been calibrated and the test is extremely repeatable. | Correct the environment requirements and associated test procedures so that this statement is true or remove the statement. |  |

Resolutions for comments 173 and 140 were presented and rejected repeatedly.

| 142 | Foegelle, Michael | 5.3.1 | 19.02 | T | A shielded room without any losses is unlikely to provide any desirable channel characteristics. It certainly cannot be calibrated through the methods described here. | Require a fully anechoic environment for this type of test. |  |
| --- | --- | --- | --- | --- | --- | --- | --- |

Reject since intent of COAT is to allow for non-flat channels with multipath if desired.

| 143 | Foegelle, Michael | 5.3.1 | 19.02 | T | The document provides no direction on determining the channel characteristics or other performance criteria of the environment. | Provide a site validation method for determining if the environment is suitable for the desired tests. |  |
| --- | --- | --- | --- | --- | --- | --- | --- |

Task Group suggested to reject as this is address in 5.3.3.2.

| 144 | Foegelle, Michael | 5.3.1 | 19.02 | T | Any antenna used as a reference for calibrating a test environment should be calibrated, not "characterized". | Require the use of a calibrated reference antenna. |  |
| --- | --- | --- | --- | --- | --- | --- | --- |

Reject comment since calibrated is a subset of Characterized.

| 145 | Foegelle, Michael | 5.3.1 | 19.02 | T | Is this text proposing that two antennas are used to connect the WLCP to the DUT in an OTA fashion? One to transmit from the WLCP to the DUT and the other to receive? | Require the use of a measurement antenna for bi-directional communication to the DUT. That antenna's performance will be determined as part of the entire path loss between the quiet zone and the WLCP or other test equipment. |
| --- | --- | --- | --- | --- | --- | --- |
| 150 | Foegelle, Michael | 5.3.2.1 | 19.16 | T | One calibrated antenna with known boresight gain must be used as a reference for determining the path loss of the environment, including the free-space path loss from the quiet zone, the gain/loss of the measurement antenna, and the cable loss to the WLCP. The measurement antenna is part of the path loss that will be measured in this process. Thus, the actual gain of the antenna is not needed. It is not used to determine the path loss, it is PART of the path loss. | Correct the terminology to apply appropriate functionality and requirements to each of the two antennas. |
| 1031 | Lemberger, Uriel | 5.3.2.1 | 19.16 | T | Transmit received reference antenna is not defined and not aligned with the naming of the antennas in the pictures in figures 6 7 and 8 | clearly define each antenna that is used in the tests and use consistence naming of each antenna. |
| 1036 | Lemberger, Uriel | 5.3.2.2 | 20.01 | T | Naming of antennas in the pictures are not consistent and not clear. | rename calibrated antenna to calibration antenna, rename Directional antenna to reference antenna. Add text that both Reference antenna and calibration antenna shall be calibrated and Reference antenna should be directional antenna. |
| 3 | Ward, Dennis | 5.3.3.1 | 20.32 | T | This step references a transmit reference antenna, but Figure 6 indicates this as a directional antenna. I believe these are the same radiating element? | Update Figure 6 to replace Directional Antenna with transmit reference antenna, or update step (b) to state directional antenna, or other mechanism to more clearly define and depict the antenna identified in this step. |
| 157 | Foegelle, Michael | 5.3.3.1 | 20.32 | MT | This text describes the measurement antenna (the one that will be used during measurements of the DUT/SUT), not the reference antenna (the one used to calibrate the path loss). | Correct the terminology. |

As suggested in 07-2551r5, change the nomenclature of the antennas to SUT, Reference, and Calibration rather than using any other argumentative nomenclature.

| 148 | Foegelle, Michael | 5.3.2.1 | 19.11 | MT | A screen room is different from a shielded room and typically does not have 95 dB of shielding effectiveness. | Remove the term "screen". "An RF shielded room, ….", "… of the shield…" |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 149 | Foegelle, Michael | 5.3.2.1 | 19.11 | T | Shielded rooms without an absorber lining are not suitable for this test environment. | Remove references to shielded rooms beyond that of the shield around an absorber lined room. |  |

Reject these comments since options are given for various channel characteristics.  Anechoic chambers are allowed if they yield the desired characteristics although absorber is not required.

| 946 | Ammann, Keith | 5.3.2.1 | 19.11 | T | An "RF shielded room" is NOT a reliable over-the-air environment. Figure 6 shows an "RF shielded screenroom" with part of it "lined with RF absorbant material". This still isn't an adequate, repeatable, over-the-air environment. Even directional antennas have sidelobes and backscatter that will cause reflections off the metallic walls. An anechoic chamber or properly constructed open air test site (OATS) would be examples of over-the-air environments that could be reasonably calibrated for repeatable results. People, equipment racks, etc. will also impact this kind of environment and reduce the repeatability. | Hopefully these issues are well understood, and people doing this sort of testing would take appropriate precautions, however, this document should mention these issues since not everybody referencing the document will have that level of experience. |
| --- | --- | --- | --- | --- | --- | --- |
| 1113 | Matta, Sudheer | 5.3.2.1 | 19.11 | T | Testing in anything other than a fully anechoic chamber is rather useless. The reflections affect the test quite a bit. | Change the first sentence to "An RF sheilded anechoic chamber" or something similar in the spirit of the comment |

Reject this comment since multipath channels can be repeatable.

| 530 | Ojard, Eric | 5.3.2.1 | 19.12 | T | start and end frequencies of "frequency band being tested" are ambiguous. does this imply regulatory bands like ISM & UNII? or does it imply the applicable channels for the PHY (e.g., clause 17 vs 19)? | Specify start and end frequencies of "frequency band being tested" |
| --- | --- | --- | --- | --- | --- | --- |
| 660 | Aldana, Carlos | 5.3.2.1 | 19.12 | T | start and end frequencies of "frequency band being tested" are ambiguous. does this imply regulatory bands like ISM & UNII? or does it imply the applicable channels for the PHY (e.g., clause 17 vs 19)? | Specify start and end frequencies of "frequency band being tested" |
| 777 | Moorti, Rajendra | 5.3.2.1 | 19.12 | T | start and end frequencies of "frequency band being tested" are ambiguous. does this imply regulatory bands like ISM & UNII? or does it imply the applicable channels for the PHY (e.g., clause 17 vs 19)? | Specify start and end frequencies of "frequency band being tested" |

Change ‘in the frequency band being tested’ to ‘over the channel frequencies being tested’.

| 532 | Ojard, Eric | 5.3.2.1 | 19.41 | T | start and end frequencies of "band being tested" are ambiguous. does this imply regulatory bands like ISM & UNII? or does it imply the applicable channels for the PHY (e.g., clause 17 vs 19)? | Specify start and end frequencies of "band being tested" |
| --- | --- | --- | --- | --- | --- | --- |
| 662 | Aldana, Carlos | 5.3.2.1 | 19.41 | T | start and end frequencies of "band being tested" are ambiguous. does this imply regulatory bands like ISM & UNII? or does it imply the applicable channels for the PHY (e.g., clause 17 vs 19)? | Specify start and end frequencies of "band being tested" |
| 779 | Moorti, Rajendra | 5.3.2.1 | 19.41 | T | start and end frequencies of "band being tested" are ambiguous. does this imply regulatory bands like ISM & UNII? or does it imply the applicable channels for the PHY (e.g., clause 17 vs 19)? | Specify start and end frequencies of "band being tested" |
| 949 | Ammann, Keith | 5.3.2.1 | 19.41 | T | On p.20 (line 35) it says to "confirm a flat channel (ripple < +/-0.5dB) in the RF channel or channels being used for the test". So, on p. 19 is the +/- 0.5dB referring to the whole 5GHz band or just across a single RF channel modulation bandwidth? In general the terms "band" and "channel" can be confusing when discussing over the air testing. "Band" could refer to the whole 5GHz band or one of the commonly defined sub-bands (e.g., 5-15-5.25GHz), or the bandwidth of the modulated signal. "Channel" can refer to the center frequency of an RF carrier (e.g. channel 36), or it could refer to the characteristics of an "RF channel" (e.g. path loss and delay spread). For example, the statement referenced above uses the term "channel" 3 times, with at least 2 different meanings. | Add definitions specifying what is meant by "Band" and "Channel". Additionally, clarify the referenced text. |

Change ‘test environment for band being tested’ to ‘test environment for channel frequencies being tested’.

| 431 | Victor, Dalton | 5.3.2.1 | 19.21 | T | Calibration antenna's characteristics are not specified | Specify that calibration antenna should be approximately unity gain. |  |
| --- | --- | --- | --- | --- | --- | --- | --- |

Reject comment since it doesn’t make sense.

| 1032 | Lemberger, Uriel | 5.3.2.1 | 19.34 | T | Turn table is not applicable in case of more than one antenna in the DUT since each antenna has different pattern and different distance from the WLCP directional antenna. Also the calibration process cannot provide a consistent method of calibration for different DUT platforms. | add text that measurement are comparable when comparing different DUT in the same platform with the same antenna system. |
| --- | --- | --- | --- | --- | --- | --- |

Reject comment since suggested resolution doesn’t match comment and commentor has not elaborated.

| 1033 | Lemberger, Uriel | 5.3.2.1 | 19.34 | T | Turn table is not applicable in case of more than one antenna in the DUT since each antenna has different pattern and different distance from the WLCP directional antenna. Also the calibration process cannot provide the same antenna gain and path loss from the WLCP antenna for all locations of the DUT antennas. | add text that the described calibration method is applicable for DUT/SUT with one antenna only |
| --- | --- | --- | --- | --- | --- | --- |

Reject comment since multiple path losses can be measured and used during measurement.

| 1034 | Lemberger, Uriel | 5.3.2.1 | 19.34 | T | Turn table is a must for testing DUT with different platforms because each device has different antenna pattern and there is no reason to prefer a specific angle of the antenna, and there is no way to compare different devices with different antenna pattern.Measuring one point has no meaning since it is not rpeatble between different DUTs. | make the usage of rotation table mandatory and copy same methodology used in OTA environments regarding the usage of the turn table except that here the rotation cannot be continuous and calibration should be repeated for each rotation angle measurement. |
| --- | --- | --- | --- | --- | --- | --- |

Reject comment since device can be re-oriented without using a turntable.

| 1035 | Lemberger, Uriel | 5.3.2.1 | 19.35 | T | Like my previous comments on 5.3.2.1, spherical positioning has no meaning for a system with more than one antenna. | either remove this option or add text that this is applicable for DUT/SUT with one antenna only |
| --- | --- | --- | --- | --- | --- | --- |

Accept comment and remove lines 19.35 and 19.36.

| 1029 | Lemberger, Uriel | 5.3.2.1 | 19.43 | T | Clause 5.2 includes table 2 with many important parameters that are essential for test. Similar information is needed also for COAT. | Add informatio in table (or other format) that includes the data in table 2 with the applicable values for the enviroment dscribed in clause 5.3 |
| --- | --- | --- | --- | --- | --- | --- |

Reject comment since the only thing that matters is the channel characteristic.  Table 2 tries to approximate a repeatable channel whereas COAT already provides a repeatable channel.

| 155 | Foegelle, Michael | 5.3.3.1 | 20.29 | T | What's the definition of "free" of extraneous noise? What level is allowed for extraneous noise? 95 dB of isolation is insufficient if there's a +50 dBm transmitter right outside the shield. Why is this requirement repeated here? Other sections of the document place a requirement on the allowed noise level. | Indicate the allowed noise level (and how it's measured) or remove this requirement since it makes no sense. |
| --- | --- | --- | --- | --- | --- | --- |

Accept comment and change “no extraneous noise present” to “no extraneous noise >-164dBm/Hz”.

| 158 | Foegelle, Michael | 5.3.3.1 | 20.32 | MT | The measurement antenna should be pointing to the DUT, especially since it's required to be directional by other parts of this section. | Re-write the bullet to indicate placement and orientation. |
| --- | --- | --- | --- | --- | --- | --- |

Accept comment and append “and pointing towards SUT” to the sentence.